FDA Inspection Report on Michigan Biologic Products Institute

The FDA inspection report on Michigan Biologic Products Institute Dated 2/20/98

[Transcriber Note]: (-WO-) equals areas the FDA whited out.

Included here are pages 02 through 12 of the report.

Begin Page 02:

1. The manufacturing process for Anthrax Vaccine is not validated.  For

example,

a.  The formulation tank has not been qualified for long term

storage of formulated bulk Anthrax.  Storage times have varied from one

week to four months between formulation and filling. Lot FAV033 was

formulated on 8/27/96, however it was not filled until 12/23/96.

b.  The formulation tank has not been qualified for mixing time,

demonstrating homogeneity of the suspension.  Mixing time is not

specified in the batch record prior to filling and during the filling

operations.  The product is (-----WO----) and settles quickly in the

tank.

c.  The firm did not perform media fill challenges to validate

aseptic manufacturing after harvest from the holding tank.  These

operations include the transfer of the sublots from building (-WO-) to

building (-WO-) for formulation.

Media fills are performed on fermentation and harvest trains, however

not on a scheduled basis.

d.  There is no validation of (--WO---) as a sporicide in anthrax

production and potency testing facilities.

e.  The analytical methods for determination of and (----WO----) in

Anthrax Vaccine are not validated with respect to accuracy, precision,

linearity, specificity and limit of detection.

f.  There is no validation of the length of time sublots are held

until they are used in a lot.  Sublots have been held longer than 3

years prior to use.  There is no stability data to support this hold

time.

Begin Page 03:

- - For example, sublot AV370 was produced and placed in coldroom

(-WO-) of building (-WO-) further referred to as (--WO--) in 2/94 until

7/97 at which time it was used to produce lot FAV040

- - Sublot AV450 was produced and placed in (--WO--) 5/95 until 3/97

at which time it was used to produce lot FAV035

- - Sublot AV456 was produced and placed in (--WO--)in 5/95 until

3/97 at which time it was transported to the formulation room of

building (-WO-) with (-WO-) other sublots to make FAV039.  Here it was

discovered that AV456 was contaminated with mold, and it was destroyed.

g.  The reference standard used for potency testing is lot FAV009,

produced March 1991.

h.  There are no expiration dates for the working spore

concentrations (virulent or avirulent strains).  For example, the

production strain, (-------WO--------) was used to produce sublot AV216

as early as 3/92 and sublot AV450 as late as 4/95.

i. (--------------WO----------------)testing for Anthrax sublots used

sublot AV462 with a (--WO--) content of 23 ppm.  The specification for

(--WO--) in Anthrax vaccine is 15-30 ppm.  There is no BF testing at

15ppm or 30ppm.

k.  Prior to August 1997, the (----WO---) filters used for harvest of

Anthrax vaccine were neither validated nor integrity tested.  This

filter is the only sterile filtration step in the Anthrax manufacturing

process.

l.  Validation of microbial retention by the (---WO----)filters used

for harvest of Anthrax vaccine was performed only with (---WO---)media,

which is used in tetanus production.  Studies were not performed using

Anthrax product or media.

** "m" is missing inspector may have failed to use that letter, no

indication of white out here**

n.  WFI used in the production of Anthrax sublots in

Begin page 04:

building (-WO-) transported from building (-WO-) in a stainless steel

tank.  There is no validation to assure that the WFI retains it's

critical quality attributes.

o.  There is no completed cleaning validation of product contact

equipment.

2.  There are no written procedures, including specifications, for the

examination, rejection, and disposition of sublots of Anthrax and

Rabies.

a.  Sublots are tested at the time of production, and are not

retested prior to formulation.  For example, sublot AV450 was produced

in 5/95 and then used in formulation of lot FAV039 in 3/97.

b.  Quarantined materials are held for extended periods.

- - For example, sublot AV216 was placed in (-W0-) on quarantine in

3/92 and was not destroyed, for low antigen content, until 5/97.

- - Sublot AV222 was placed in (--WO-) in 4/92 and was removed and

destroyed in 5/97 due to mold.

- - Sublot AV493 was manufactured in 8/96 and is still in quarantine

in (--WO--) for low antigen content.

3.  Potency testing of Anthrax Vaccine requires either testing 1

finished product vial, an aliguot from the formulated bulk tank, or a

pilot bulk sample.  There is no data demonstrating that these samples

are representative of the lot.

In addition, expiration dates are assigned based on the latest valid

potency test.  There is no correlation between this date and formulation

of bulk or filling of the finished product.

Begin page 05:

4.  There is no written justification for redating lots of Anthrax

vaccine that have expired.  "Redating" testing consists only of a

potency test.  There is no documentation of testing for

container/closure integrity or container/closure compatibility for

periods up to 7 years.  In addition there is no analytical testing

identifying and demonstrating the absence of degradants.

There is no written SOP for redating, including when redating will be

performed in order to extend the expiration period.

- -Lot FAV023 was tested for redating 2 times in 1997 and failed.  It

also failed twice on stability in 1997.  It is scheduled to be retested

for redating on 4/21/98.

Anthrax lots that are submitted for redating are released by CBER with

alternate lot numbers to indicate the redate.  However product is not

labeled with the alternate lot number.

- -Lot FAV020 (initial date of potency 4/13/93) was submitted for

redating as FAV020-1 and labeled on 2/6/98 as FAV020

For Anthrax Vaccine lots #FAV008 through #FAV016, the firm unpacked the

vials from the cartons and removed the labels (the labels were removed

by soaking in alcohol).  The firm does not have a written procedure for

performing unpackaging of vials and removal of labels.  Also, the firm

does not have documentation of performing reconciliation of the vials

before and after this operation.

5.  Regarding the firm's stability program for Anthrax:

a  The firm's stability program did not start until 1997.  Stability

testing consists only of performing release tests at various intervals

and does not address product degradation.  There is no justification for

placing lots manufactured as early as 1991 into the stability program.

Begin page 06:

b.  The firm does not have a system in place to investigate and

report stability failures.  (All results noted below are from samples

stored at 2 - 8 degrees C.  Accelerated stability test results are not

included.) For example,

- - Lot FAV010 was filled on 7/1/91 and submitted for redating on

10/11/94, having passed all potency testing.  It was placed in the

stability program (zero time) on 10/7/97 and tested for potency. The

"zero time" potency for this lot is recorded as "unsatisfactory".  There

is no investigation into this result nor is there any additional potency

testing.

- - Lot FAV011 was filled on 10/17/91 and submitted for redating on

10/11/94, having passed all potency testing.  It was placed in the

stability program (zero time) on 10/27/97, and tested for potency.  The

"zero time" potency for this lot is recorded as "No Test" having not met

the (-WO-) dilution criteria.  Test records indicate it was tested for

potency 11/24/97.  There is no investigation into this result nor is

there any additional potency testing.  In addition, the test result for

(---------------WO------------------), dated 10/30/97 is 11 ppm.

(Specification is 15-30ppm.)  This lot was not rejected nor placed in

quarantine.

- - Lot FAV013 was filled on 10/25/91 and submitted for redating on

11/22/94, having passed all potency testing.  It was placed in the

stability program (zero time) on 11/12/97, and is recorded as "No test",

having an invalid test.  Test records indicate it was tested for potency

on 12/5/97. There is no investigation into this result nor is there any

additional potency testing.

- - Lot FAV18 was filled on 7/28/92 and submitted for redating on

6/11/96, having passed all potency testing.  It was placed in the

stability program (zero time)  on 6/10/97, and is recorded as

"Unsatisfactory".  Test results dated 7/7/97 indicate it failed potency

specifications.  There is no investigation into

Begin page 07:

this result nor is there any additional potency testing

- - Lot FAV022 was filled on 2/9/93 and submitted for redating on

10/15/96, having passed all potency testing.  It was placed in the

stability program (zero time) on 10/7/97, and is recorded as having an

"Unsatisfactory valid test".  Test results dated 10/31/97 indicate it

failed potency specifications.  There is no investigation into this

result nor is there any additional potency testing.

- - Lot FAV023 was filled on 12/13/93 and passed a potency test on

3/29/94.  It was submitted for redating on 4/2/97 and was placed in the

stability program (zero time) at the same time.  It is reported as

failing potency on 4/2/97.  It was tested again on 8/12/97 and is

reported as failing potency.  A fourth potency test conducted on 10/6/97

is listed as passing by 0.01.  There is no investigation into the

original result and justifying the additional testing.

- - Lot FAV040 was filled on 11/13/97 and placed in the stability

program (zero time) on 11/19/97.  It is reported as having an "invalid"

potency test on 11/19/97.  There is no investigation into this result

nor is there any additional potency testing.

c.  The firm's SOP(s) for handling manufacturing

deviations/departures does not address when a lot should be monitored on

stability.

6.  There has been no investigation into numerous "invalid" potency test

results for lots.  For example:

- - Lot FAV021 was filled on 11/24/92, having a passing potency test.

It was tested again on 10/15/96 and failed potency.  It was tested again

on 1/28/97 for redating and passed.  There is no investigation into the

test failure nor justification for

Begin page 08:

retesting the lot.

- -Lot FAV025 was filled on 4/24/94, having a passing potency test.  It

was tested again on 4/22/97 and failed potency testing.  There is no

investigation into the test failure.

- -Lot FAV028 was filled on 6/2/95.  It was not tested for potency

until 7/9/96 when it failed the test. It was tested again on 8/27/96 and

passed. There is no investigation into the test failure nor

justification for retesting the lot.

- - Lot FAV041 was filed on 11/18/97.  It had an "invalid" potency

test on 9/30/97.  There is no investigation into this invalid test.

- - Lot FAV042 was filled on 11/21/97.  It had an "invalid" potency

test on 10/29/97.

- - Lot FAV043 was filled on 12/25/97.  It had an "invalid" potency

test on 11/18/97.  There is no investigation into this invalid test.

- - Lot FAV044 was filled on 1/7/98.  It had an "invalid" potency test

on 12/8/97.  There is no investigation into this invalid test.

7.  The firm's SOP (---WO--- 36 spaces of WO!----------)dated 9/3/96,

requires that vials discarded as rejects be counted, however, it does

not specify limits for when a lot should be investigated or rejected as

a result of this lot reconciliation.  For example:

-- Lot FAV016 had 6579 vials rejected due to particulates

during post filling inspection.  These particulates were not

identified,  nor was an investigation conducted.  The batch was

released.

Begin page 09:

- - Lots FAV028, FAV031, FAV033, and FAV038 had 3323, 2441, 2509,

1347 vials rejected respectively for low volume during post filling

inspection.  There was no investigation conducted.

- - Lot FAV035 had 409 vials rejected for faulty closure during

post filling inspection.  There was no investigation conducted.

8.  The firm does not have specifications for time limits at which the

product can be exposed to room temperature conditions during filling,

labeling and packaging operations (repeat observation).

- - Lot PAV036 was at room temperature for (-WO-) hours and then

the filling operation was aborted, it was placed back in the

refrigerator (deviation report #97DAV34).

In addition, there is no stability information regarding product

exposure to room temperature.  Prior to 1996 the firm did not monitor

the length of time at which the product was exposed to room temperature

conditions during the filling operations (FAV009-FAV015).

9.  The firm's procedures for Environmental Monitoring of critical

production areas do not require that additional cleaning and increased

sampling be performed when environmental action limits are exceeded.

When environmental monitoring action limits are exceeded during filling,

investigations do not consider environmental monitoring results during

production of sublots, sterility test results of sublots and sterility

test results of final product.  In addition, when a sterility retest is

performed during stability testing, no investigation is performed.  For

example:

- - Lot FAV029 was filled on 8/11/95 and passed sterility testing.

On 9/23/97, during stability testing (--WO?)

Begin page 10:

it required a sterility retest.  The contaminant was identified as

Penicillium species.  The product passed a retest.  Production records

for sublots used to produce FAV029 indicate (----WO----) were bulked to

produce the lot.  Sublot AV383 had an initial sterility failure on

5/17/94 (Rhodocoddus species); Sublot AV390 had an initial sterility

failure on 7/19/94 (Propionibacterium acnes).  Both sublots passed

sterility retest.  There is no environmental monitoring data from

preparation of the sublots. On 8/11/95, during filling of lot FAV029,

environmental monitoring testing found the following on critical

surfaces:  Cladosporium species, Alternaria species, Micrococcus

species.  Bacillus subtilis, Staphylococcus saprophyticus,

Staphylococcus epidermidis, and Staphylococcus capitis.

- - Lot FAV032 was filled on 10/26/95.  On 7/28/97, during stability

testing (-WO-) it required a sterility retest.  The contaminant was

identified as Penicillium species.  The product passes a retest.  The

lot was formulated on 9/21/95.  Two operators performing the formulation

exceeded action limits on viable monitoring.  Four CFU were sampled from

one of the operator's gloves and identified as Penicillium species.

- - Lot FAV035 was filled on 2/5/97.  On 8/11/97, during stability

testing (--WO--) it required a sterility retest.  The contaminant was

identified as Bacilius cereus.  The product passed a retest.  The lot

was formulated on 1/9/97.  Environmental monitoring exceeded action

limits in the gowning area prior to formulation identifying the

following:  Staphylococcus capitis, Micrococcus species, Bacillus

coagulans, and Corynepacterium species.  In addition, photohelic gauges

were out of range during this time indicating insufficient air pressure

in critical areas.

--The firm does not trend multiple contaminations with

microorganisms in sublots.  For example, between 4/94 and 2/95,

(----WO---) were produced of which 23 were discarded due to some kind of

microbial contamination.  Lots FAV029, FAV030 and FAV031

Begin page 11:

were whole or partially formulated from those sublots not discarded in

this time period.  In January and February 1997 of 12 consecutive

sublots produced, 5 were discarded for microbial contamination.  The

others were included in lot FAV039.  In September and October 1997, of

(---WO---) produced, 6 had contamination and two of those were retested

and released for formulation.  The remaining sublots were formulated

into FAV045 or FAV046.

10.  Recording of data in building (-WO-) from room (--WO--) log books

in room (-WO-) is accomplished by viewing the results through the UV

pass box and is not checked for accuracy prior to discarding the

original data.

11.  Specifications for the release of sublots were not formally

established until 1995.

12.  The firm does not have a current SOP for environmental monitoring

in the Anthrax production facility.  The firm has replaced it's previous

environmental monitoring SOP with a centralized procedure that

references area specific monitoring plans.  However, the Anthrax

specific plan has not been finalized.

Anthrax Building Facilities Conditions:

13.  In room(s??) (------WO--------) of the Anthrax production facility,

we observed peeling paint, exposed duct and pipe work, insulation

peeling off the pipes, and rusty steam and gas lines.

14.  Compressed air used to perform positive pressure transfers of sterile

products (Anthrax and Rabies) is central plant air and ?? is not

monitored.  The (-WO-)micron filters used at the point of use are not

integrity tested.

15. Rooms (-------WO------) are not environmentally controlled.  There

was no active environmental monitoring of aseptic

Begin page 12:

manufacturing activities until 1996.  During the 1996 production of

sublots, more than half of the sublots had environmental monitoring

excursions.  There was no tracking of these events and no significant

corrective action taken until 10/96.

16.  Plant steam is used for sterilizing production vessels and

glassware in buildings (---WO----) and is not monitored or controlled.

17. Poor facility arrangements exist for aseptic processes in building

(-WO-) room (-WO-) in that media is made, dishes washed, equipment and

glassware autoclaved, as well as the production processes of

fermentation, inoculation, and harvest all occur in this one room

simultaneously.

18.  Regarding cold storage of critical seed stock:

a.  In the Anthrax production suite the logs for

refrigerator/freezer(--------------------------------

--------WO---------------------------------) are incomplete.  The logs

do not match the refrigerator/freezer contents.  The Anthrax

refrigerator/freezer contained unlabeled vials.

b.  There is no segregation of the master spore concentrations and

the working spore concentrations of both the virulent and old (---WO---)

strains in (----WO---) Anthrax production and potency testing

facilities.

c.  The keys for all refrigerator/freezers in building

(-----WO------) and building (-WO-) were found on top.

19.  There is no SOP for change over in building (-WO-) Anthrax

Biosafety Cabinets (BSC).  Both inoculum preparations and aseptic sublot

formulation occur in these hoods.

Anthrax vaccines expired

May 25, 1998

By Rod Hafemeister

Belleville News-Democrat

The U.S. Navy injected sailors  in the Persian  Gulf  with  a  5-year-old batch of anthrax vaccine, two months after  federal  regulators said  the vaccine had been given a new expiration date improperly. That's according to U.S. Food and Drug Administration  records  that  show the state-owned  Michigan Biologic Products Institute  relabeled  a 1993 batch of  anthrax  vaccine in February to extend its shelf life  past 1996. Sailors  on  two  aircraft carriers in the gulf say medical  records list  that same  batch  - FAV020 - as  the  one  given to  the  sailors in  April, according  to  interviews. There is no indication the vaccine could harm the sailors, but the FDA can't certify that it's safe either.  Changing the date of a vaccine is legal, FDA spokesman Lenora Gelb said, but the product has to be recertified as  safe  and  effective.  Batch FAV020 was tested for potency, but not for other problems, FDA records show.

The Michigan plant made all of the  more  than  seven  million  doses  of anthrax vaccine the Pentagon said it will use  this summer  to begin  its program to vaccinate  every  active  duty,  reserve  and  National  Guard service member - 2.4 million in all.

But FDA records show that since 1993 the Michigan  plant  repeatedly  has been cited for violations of the FDA's Good Manufacturing Practices rules and other quality control problems. The records were obtained through the Freedom  of  Information  Act  and  with  the  help  of  GulfWatch, a  veterans' advocacy organization based in Hannibal, Mo.

The records include a March 11, 1997, letter from the FDA threatening to revoke the plant's licenses if it did not correct its problems.  In February, FDA inspectors spent two weeks in the plant, with  much of

their attention focused on the anthrax vaccine line, and concluded:

" The manufacturing process for anthrax  vaccine is  not  validated."     Among the problems:

" Vaccine was being stored for years at a time with no way of telling

whether the vaccine remained effective and uncontaminated.

* Different  batches  of  vaccine with  widely varying  manufacturing

dates were mixed together to create a final batch.

* Until August 1997, filters  used to  harvest the  vaccine were  not

validated as effective nor was their integrity tested.

* No written justification existed for redating batches of  vaccine

that had expired, and there was no proof the plant had even  tested

the batches for contamination  before  giving them new  expiration

dates.  Batches that failed potency tests were sometimes retested

until they passed, with no explanation of why they failed.

Run by the Michigan Department of Community Health, Michigan Biologic Products Institute is the only licensed manufacturer of anthrax vaccine for humans in the United States. For decades, the plant  made relatively small amounts of anthrax vaccine for veterinarians and other workers who could be exposed to the deadly livestock disease.

But the vaccine has  gained  new  attention  in  recent  years,  in  part because Iraq's Saddam Hussein had an  active biological  warfare  program and in part because  anthrax  is  believed  to  be  one  of  the easiest biological weapons for terrorists to make.

"We know there have been problems  and  we  are  working  on  them," said Geralyn Lasher,  spokesman  for  the  Michigan  Health  Department. "But remember, the anthrax vaccine has more testing done on it than pediatric vaccines."

Vaccine batches submitted to FDA for new  expiration dates  are  supposed to be relabeled with an alternate Lot Number  that indicates it  has been redated. Inspectors found that  didn't happen,  specifically  identifying Lot FAV020, which was originally dated  as being  tested for  potency  on April 13, 1993, and redated on Feb. 6, 1998.

Sailors on both the USS Independence and the USS Stennis in  the  Persian Gulf in April received vaccinations  from Lot  FAV020, according to three sailors on the ships who say that lot number is in their medical records. The sailors were reached through Lori Greenleaf of Morrison, Colo., whose son Eric Julius is stationed  on  the  USS  Independence.  One  of  those  sailors reached aboard the Independence was Nhut M.  Nguyen,  who refused  the shots but claims other sailors' medical records show the questionable  batch number.  An estimated 14 sailors on the two  ships refused  to take  the shots and  have  received  administrative  punishments  ranging  from  reductions in rank  to discharge  from the Navy.  Navy  spokesmen did not return repeated calls last week.  Pentagon spokesmen  referred questions

to the Navy.  The Pentagon in March ordered anthrax vaccinations for all troops in the Persian Gulf, including the sailors on the Independence and the Stennis.

After its February inspection, the FDA didn't close  the  Michigan  plant

- they didn't get the chance.  The anthrax  line  was shut  down in March to begin an 18-month, $20 million  renovation  program  paid for  by  the Army. But in an April 7 letter, John Taylor III, FDA's senior advisor for regulatory operations, warned that Michigan  Biologic Products  needs  to make "significant improvements" before resuming production. Critics think  the FDA buckled to pressure from the  Pentagon and  let  the Michigan lab off easy.

"The FDA has really been talking out of both  sides of  their mouth – the standards were sloppy but  the  product  was fine," said  Lingg Brewer, a Michigan state lawmaker who has been investigating  the lab.  Meryl Nass, one of the few nonmilitary doctors who has studied  anthrax, said she was appalled by the FDA  reports. "It  seems  that  thousands  of  vials were discarded if there was visible evidence of contamination  or  development of  particulates," she  said. "But  the  remaining  vials  in  those lots (batches) were released for use in people without further testing."

Pentagon officials maintain the anthrax is safe and effective. On Friday, Defense Secretary William Cohen announced that  the Pentagon will proceed this summer with its plans to  vaccinate 2.4 million active duty, reserve and National Guard service members over the next six to seven years.  The same day, President Clinton called for stockpiling up to 30 million doses of vaccine for defense against biological terrorism.

Attorneys for Veterans for Integrity in Government, a  Washington,  D.C., advocacy group,  have been  attempting  to stop  the anthrax  vaccination program, arguing that the Pentagon cannot show  that it will either work against anthrax used as a weapon or that it is safe in the long term.

(c) 1998 Belleville News-Democrat

This article was published here with permission from Rod Hafemeister.

Scientists describe anthrax vaccine as an outdated 'disgusting mix'

May 25, 1998

By Rod Hafemeister

Belleville News-Democrat

Researchers have described the anthrax vaccine the  Pentagon  intends  to

give millions  of  troops  as  "1950s  technology  unimpeded  by  medical

progress." Critics have called it a "disgusting mix" and a "soup" because

the manufacturing process leaves often unknown  compounds in the vaccine.

It is made by growing a weakened strain of anthrax  bacteria, killing the

bacteria and  filtering out  a protein called protective antigen, or P.A.

Researchers have found that  P.A. causes the body to develop a resistance

to the lethal toxins produced  by anthrax.  One  criticism  is that there

is a wide variation from batch to batch in the  amount of P.A., and  that

no standard exists for the amount of P.A. needed to produce immunity.

Researchers don't even know if the  current 18-month  series of six shots

is necessary - it is based on a  single study  done in the  1950s.  Chuck

Dacey, a spokesman for  the Army  Research and  Material  Command at Fort

Detrick in Maryland, said there is an ongoing clinical trial to determine

whether fewer doses can be used.

Pentagon officials have  repeatedly  expressed  full  confidence  in  the

effectiveness and safety of the vaccine. But for more than a decade, they

also have been funding a series of attempts to  develop a better vaccine.

Last year, a highly purified vaccine was developed  at Fort Detrick  that

shows promise but is  several  years  away  from  production.   "The next

generation  vaccine is  still  very  much in  research  and development,"

Dacey said.

(c) 1998 Belleville News-Democrat Inc. All rights reserved.

This article was published here with permission from Rod Hafemeister.